TDS8200, CSA8200 Technical Report

Date: June 27, 2004

Subject:	Memory Erasure for TDS8200 and CSA8200 Oscilloscopes

These products contain an open architecture PC with REMOVABLE hard drive.  All user storable data – waveforms and instrument settings – are stored on the removable hard drive.  The hard drive is located on the rear panel.  Remove the hard drive for storage or destruction.  Additional hard drives can be purchased from Tektronix.

Alternately, there are DOD approved scrubbing software packages available for the hard drive.  Tektronix has no recommendations regarding the available packages.  After the hard drive has been scrubbed, the operating system and oscilloscope software can be re-installed using the Operating System Restore and Product Software CDs that came with the Oscilloscope.

The PC motherboard BIOS and CMOS settings can be reset to their factory default by following steps outlined in the Operating System Restore CD instructions.  Choose the menu option “Update motherboard BIOS firmware and CMOS settings” when presented during the OS restore sequence.

Resetting the motherboard BIOS and CMOS settings, reinstalling the Operating System, or reinstalling the Product Software will not affect calibration of the instrument because all mainframe related calibration constants are stored in non-volatile memory on the acquisition board within the mainframe rather than on the hard drive.  This allows complete erasure/removal of any secure data without affecting oscilloscope calibration.  It also allows the oscilloscope to be calibrated in a non-secure site then used in a secure area without need for recalibration.

Floppy disk drives are standard on all TDS8200 and CSA8200 models.  Remove all floppy disks and format them (don’t use “Quick Format”), store them, or destroy them.

Read-only CD-ROM drives are standard on all TDS8200 and CSA8200 models.  Remove all CDs.  Rewritable CDs can be formatted, stored, or destroyed.  Non-rewritable CDs can either be stored or destroyed.

Turn power off for at least 20 seconds to clear volatile memory.

If you have any questions, contact the Tektronix Technical Support Center at 1-800-833-9200*3.